Mechanical Engineering Department

ME 372 Measurement Laboratory (METROLOGY)

Course Instructor: Prof. Ramesh K Singh

Introduction

The requirements of industries for parts with greater precision, tighter tolerances and higher reliability have accentuated the developments in the science of production and inspection. Quality production cannot be realized without sophisticated measurement and inspection techniques.

Metrology essentially deals with the instruments and techniques for accurate inspection of components for attributes like dimensions, form, surface finish, etc.

The present laboratory course on Metrology is designed to cover basic principles for the dimension and form measurement and the associated instruments and gages. It comprises of the following broad topics:-

Course Structure

Principles of Engineering Measurements.

Measurement of lengths and angles (Hand and Precision Instruments).

Limits and Fits.

Geometric Tolerances for form, position and location.

Design of special purpose inspection gages/comparators.

Inspection of complex shapes like screw threads, worms and gears.

Surface finish and its measurement.

Introduction to Modern Techniques of inspection based on the above topics.

Accordingly, following five experiments have been planned –

Experiments

- Measurement of Flatness Error/Optical Flat

- Calibration of Dial Indicator/White light interferometer

- Inspection of Parts for Geometric Characteristics/Auto-Collimator

- Inspection of Threads/Inspection of Gears

- CMM demo/Optical microscope

Mode of Assignment (Metrology)

|  | Report and viva voce | End semester test | Total |
| --- | --- | --- | --- |
| Percentage | 70% | 30% | 100% |

Text Books

- R. K. Jain, ‘Engineering Metrology’, Khanna Publishers, 1999.

- J. F. W. Gayler and C.R. Shotbolt, ‘Metrology for Engineers’ Cassel London, 1964.

- K. J. Hume, ‘Engineering Metrology’ MC Donald London, 1963.

- G. G. Thomas, ‘Engineering Metrology’, Butterworth’s, 1974.

Reference Books

- ASTME : Handbook of Industrial Metrology, Prentice Hall, 1972.

- F. T. Farago, ‘Handbook of Dimensional Measurements’, 2nd Ed. Industrial Press Inc. 1982.

General Instructions

- Study the write-up on experiments and tutorial given to you before coming for experiments.

- Maintain a separate notebook for Metrology Laboratory experiments.

- Start an experiment with a neat sketch of the component to be inspected.

- Chalk out a plan for the measurements to be carried out based on the write-up given to you. Contact the instructor concerned for further procedure.

- Do not meddle or operate any equipment in the laboratory without proper knowledge of operating them. Contact the instructor concerned to know their operation.

- Bring the calculators and other drawing instruments, wherever necessary.

- Leave your bags and other belongings on the table provided in the laboratory, as soon as you enter the laboratory.

- You will be allowed to repeat the experiments for valid absence only.

Precautions While Handling Metrology Instruments

- Do not try to operate micrometers unless the lock nut is slightly released.

- While handling the bevel protractor, do not release the lock nut completely.

- While focusing an object by NIKON microscope, initially use coarse knob until you get good image of the specimen and fine focus be used only to get a sharp image. The fine focus knob should be used only when necessary.

- Do not touch with bare hands the anvils of micrometer, measuring faces of slip gauges, eyepiece lenses of NIKON MICROSCOPE and measuring lengths of standard measuring wires. Use special cloth for handling these.

- Do not meddle with the level adjustment screw of the surface plate since it is usually set for specific levels.

- Make sure that the component to be inspected is sufficiently clean before starting measurements on it.

Submission of Laboratory Journal

- At least one observation notebook per batch should be maintained. Make sure to get it

signed by the concerned instructor and your attendance is marked.

- Submit one journal per batch in the standard institute files when you come to the Metrology laboratory for your next turn of experiments. Your names, roll numbers, batch Numbers should be clearly mentioned on it.

- For the first turn submit your write up in files. For the subsequent turns, submit the sheets duly stapled which will be filled in respective files.

Experiment No. 1

MEASUREMENT OF FLATNESS ERROR

Part A: Inspection of Surface Plate for Flatness Error

    - Theory

Surface plate is used as a datum plane for accurate measurement of dimensions of a component. It is also used for testing the flatness of other surfaces. Hence, it is necessary to inspect the surface plate itself for its flatness error.

    - Aim of Experiment

To determine the flatness error of a surface plate with the help of a sensible spirit level.

1.1.3	Procedure

- Mark the boundary lines and generators on the surface plate (as shown in Figure 1). Lengths of the lines should be integer multiples of a common reading span.

- Choose a reference plane, say ABD, from which the flatness deviations of all the points will be determined.

- Place the spirit level on the first span A-1 and note down the deviation of point, with respect to A.

- Repeat the readings for 1-2, 2-3 etc. until B.

- Repeat the same procedure for all the lines keeping in mind that all the readings independently for all the lines.

- Correct all the ends of the lines of the reference plan to zero. Correct the other lines with respect to the three points on the reference plan.

- Thus the deviations of all the points relative to the reference plane chosen are known.

- The minimum distance between a pair of parallel planes that contains all the points chosen on the surface will give the flatness error of the surface.

Figure 1. Layout on surface plate for flatness measurement.

1.1.4	Results

- Tabulate the results independently for all the lines

- Determine the flatness error

1.1.5 Conclusions

- Comment on the sources of errors and the assumption made.

- Discuss other methods of measuring the flatness of the surface plate.

Part B: Flatness Testing using Interferometry

    - Theory

The essential equipment for light wave interference is a monochromatic light source and a set of optical flats. An optical flat is a circular piece of optical glass having its two plane faces flat and parallel, and finished to the optical degree of flatness. When an optical flat is placed upon another flat reflecting surface (without pressure), it will not form an intimate contact but will lie at some angle making an inclined place, see Figure 2. If the optical flat is illuminated by a monochromatic source of light, and the eye is placed in a proper position, a number of bands will be observed. These bands are produced by the interference of light rays reflected from the top surface of the surface being inspected and lower surface of the optical flat. If path length of these two light rays differ by an odd number of half wavelength, then the condition of complete interference is achieved and a straight dark line is (fringe) is seen passing through the point C. Similarly such fringes will be seen at 3/2, 5/2, (at point H), etc. At the intermediate points, between C & H, the path difference will be even multiple of half wavelength and two components of light rays will be in phase producing a light band.

Figure 2. Principle of interferometry.

Thus, for a perfectly flat surface, alternate bright and dark bands will be seen. Any deviation from this will be a measure of the error in the flatness of the surface.

1.2.2	Aim of The Experiment

To study flatness using optical flats and monochromatic light source

1.2.3	Procedure

Take five slip gauges of workshop grade, one gauge of inspection grade, a micrometer and an optical flat. Clean all the objects sufficiently, else fringes will not be visible.

Switch ON the monochromatic light and place the object to be inspected on the stage under monochromatic light.

Position the optical flat on the surface to be inspected. Ensure that while placing or removing the optical flat from the surface, it will not slide. As this will scratch the optical flat reducing its useful life. Observes the fringes.

Repeat the procedure for all slip gauges and micrometer spindle.

1.2.4	Results

Sketch the fringe patterns for each of the objects and draw your conclusions

1.2.5	Conclusions

Compare the measurement of flatness error by surface plate and optical flats.

Experiment No. 2

CALIBRATION OF DIAL INDIACTOR

  - Theory

In most of the measurements, the dimensions to be measured are compared with a suitable reference standard. Such measurements involve the use of an instrument called comparator; which measures the deviations of a dimension from the standard reference value. The comparators however, may sometime indicate faulty readings due to their inherent manufacturing errors or due to the wear of their parts during service. The comparator, therefore, needs to be calibrated periodically.

  - Aim Of The Experiment

To calibrate the dial indicator with the help of a Dial Tester

  - Procedure

- Study the dial indicator provided. Chalk out a plan for calibration. (Range to be calibrated, intervals of the readings to be taken, reference standard to be used etc.)

- The dial indicator is fixed on the dial tester with the plunger just touching the anvil of the micrometer. The pointer on the dial indicator should be adjusted such that the zero coincides the bezel. The circular disc of the micrometer is adjusted such that the zero coincides with the cross wire of the plastic indicator.

- Rotate the thimble of the micrometer in clockwise direction, Note down the readings of the graduated disc of the micrometer against the cross wire for each interval of movement of the dial gage.

- Repeat the procedure for the range of the indicator, both in the forward and backward directions.

- Since the Dial Tester is used as reference standard. It should be justified. Choose a proper range of the dial tester and check against the slip gauges.

  - Results

- Tabulate the results as divisions on the dial gage vs. the deviations from the actual readings note from the dial tester.

- Plot a graph of calibration with the dial indicator readings vs. the error for various turns

- Tabulate the results for justification of dial tester.

  - Conclusions

- Study the calibration graph and comment about sources of error in the dial indicator under test and its measurement capability.

- Comment about the capability of the Dial tester for its use as reference standard for calibration.

- Discuss about the sources of error in the measurement and precaution to be taken.

Experiment No. 3

INSPECTION OF PARTS FOR GEOMTRIC CHARATERISTICS

  - Theory

The measurement of geometrical characteristics in modern engineering has become a factor of vital importance in the process of production due to higher and higher requirements for the accuracy of machined surfaces. For the satisfactory functioning of parts, it is not sufficient to accept or reject a part based on the size tolerances only. Geometric characteristics like errors in form, orientation, and position/location, etc. are also important for manufacture inspection and assembly.

3.2.1 Part A: Inspection of Geometrical Characteristics using dial indicator

Aim: To inspect the geometrical characteristics of the given component, viz. Circularity, Cylindricity, Concentricity, Parallelism and Perpendicularity.

Procedure

- Study the component provided.

- Identify the geometric characteristics of form, orientation, location/position etc.

- Choose a proper method and instruments and inspect the component.

# Results

- Present the measured value in tabular form independently for all the characteristics measured.

- Present the final results with a neat sketch using geometric tolerance symbols and showing the tolerance zone.

# Conclusions

- Mention which geometric characteristics could not be inspected; suggest suitable methods for measuring them.

- Discuss the sources of errors in the measurement.

3.2.2 Part B: Inspection of Geometrical Characteristics using Auto-Collimator

Aim: To inspect straightness of a machine bed using Auto-Collimator

Procedure: Read the instruction manual of the auto-collimator carefully and familiarize yourself with the principle and procedure of measurement. Position the reflector of the auto-collimator at various places on the machine bed and perform the measurements as instructed.

Results: Tabulate the results as per the instructions in the study material to evaluate the straightness of the machine bed.

Conclusions: Discuss the sources of errors in the measurement.

Experiment No. 4 (a)

INSPECTION OF SCREW THREADS

4.1 Theory

A screw thread is generally used to transmit force and motion (e.g. lead screw of a lathe, screw jack etc.) or to act as a fastener. Various types of thread forms have been developed for different applications. It is important that the machining parts of a thread pair must be produced with matching dimensions that should be closely maintained during manufacture. Hence the inspection of threads is vital for the proper functioning of assembled parts.

4.2 Aim of The Experiment

To inspect the screw thread for pitch, thread angle, major and minor diameters, pitch diameter and to determine the pitch errors.

4.3 Procedure

Part A: Measurement of Pitch, Thread angle, Major and Minor Diameters

- All these parameters can be measured using tool makers microscope (Two Coordinate Measuring machine).

- Study the manual of the microscope.

- Select proper keys on the data processor for a particular measurement and measure the parameters for at least 10 threads.

PART B: Measurement of Pitch Diameter using Three-wire Method

- The three-wire method involves taking measurement over three wires of known and identical diameters, lying on the opposite sides of corresponding grooves of thread.

- First calculate the ‘Best Size Wire’ diameter using the equation

d = P/2 sec (Q/2) where P= Pitch, Q= Included angle of threads.

- Choose the appropriate wire set from the sets of wires available and place them on the opposite sides of corresponding grooves of thread.

- Measure the dimension over the wires M, using a flange micrometer. Repeat the measurement for at least three threads.

- The pitch diameter can be calculated using the appropriate equation involving M, P, Q and d.

4.4 Results

- Tabulate all the parameters measured.

- Calculate the cumulative pitch error.

- Present the results with a neat sketch.

4.5 Conclusions

- Comments on the sources of error in the measurements.

- Mention the methods of thread inspection in mass production.

- Give your conclusion about the type of thread you have inspected.

Experiment No. 4 (b)

INSPECTION OF GEARS

5.1 Theory

Gears are mainly used to transmission of power and motion. In order to the rotary motion of the driven shaft be perfectly uniform relative to the driving shaft. It is essential that both the gears be of correct geometrical form and accurately mounted on shaft running in good quality bearings. Hence the precision with which gears are manufactured and inspected plays a vital role.

5.2 Aim of The Experiment

To inspect worm and spur gears for their parameters specified.

5.3 Procedure

Part A: To inspect worm for its pitch and profile angle.

The worm is inspected with the help of the worm and hob testing machine. Study the machine carefully, understand the movements and measure the parameters. The formula for profile angle calculation is given on the machine head stock itself.

Part B: To inspect the spur gear for its tooth thickness and angular pitch.

- Tooth thickness of spur gear is measured using gear tooth vernier caliper.

- The theoretical values of chordal width and depth of spur gear are given by

W= Nm sin (90/N)

d = Nm/2 [1+ (2/N) - cos (90/N)] where, N = No. of teeth, m =module.

- Set the chordal depth value in the vertical scale of the vernier and measure the tooth thickness by adjusting the jaws of the horizontal scale to be tangent to the tooth.

- Take readings over the entire circumference so that tooth thickness error can be calculated.

- Angular pitch is measured using tangent micrometer. Let the measurements over x teeth and (x + y) teeth M1 and M2 respectively.

Base pitch = (M2- M1) / y

- Knowing base pitch, angular pitch can be calculated. Repeat the readings at three locations so that angular pitch error can be estimated.

5.4 Results

Present the results in tabular form and with neat sketches.

5.5 Conclusions

Discuss the source of errors in the experiment.

Experiment No. 5

CMM/Measuring Microscope

- Use the microscope for getting the data

- Draw a solid model of the objects being measured.